## System Calibration

# DC: Longitudinal Regression Factor Calibration Worksheet

### Sample: Palladium

### Temperature: 298 K => Pd = 5.25*10-6 emu/Gauss/gramm (see datasheet in manual)

### DC-Parameters: 4cm, 24 Pts, 3 scans

Sequence to measure Pd:

Scan Field, oscillate, high Res enabled, 1 T to Maximum Field, 1 T steps

Measure DC

End Scan

Note old Longitudinal SQUID factor: Cold =

Note weight of Pd: mPd =

Calculate slope a and offset b of Magnetization values Mi = a * Hi + b:

( for calculation by hand: a = ((Hi-Hmean)(Mi-Mmean)) / (Hi-Hmean)2, b = Mmean – a * Hmean )

Calculate new Longitudinal SQUID factor: Cnew = Cold * a / Pd / m =

Adjust Longitudinal SQUID calibration Factor only if difference from old value > 1%

Set Field H = 1T (oscillate mode) and calculate theoretical value:

Mtheo = a * H + b

Measure Palladium and calculate difference: (Mtheo / M1T – 1 ) * 100 < 1% ?

If bigger than 1 % the procedure has to be repeated.

# AC Regression Factor Worksheet

Sample: Dysprosium Oxide (Dy2O3), Temperature: 298 K

AC-Parameters: Drive Amplitude = 3 Gauss, Frequency = 20 Hz, Blocks to Average = 20, Measurements to Average = 3, Longitude Range = auto range, Null amplitude = 0.00001

Step 1:  Set H1 = 10000 Gauss DC field.

Record old AC Regression Factor: _________________

Measure DC (4cm, 5scans, 24 points): m1 = ________________

Measure AC and compute the susceptibility AC:

m’ = ________________

AC =  m’/3 (EMU/gauss) = _________________

Step 2: Set H2 = 20000 Gauss DC field.

Measure DC moment:	m2 = ________________

Calculate DC susceptibility: DC = (m1-m2)/(H1-H2)

DC = 	________________

Step 3:

Calculate New Factor and Enter (Diagnostics/Cal. Factors/AC Cal. Menu).

New factor = (old factor) * (DC/AC) = ______________

Verification:

Measure AC moment, compute the susceptibility, and compare with the original DC.  The values should agree within .5%.  Print out the AC waveform graph of the measurement and attach to this worksheet.

m’ = ________________

ACVER = m’/3 (EMU/gauss) =________________

Calculate % Difference =  {(ACVER/DC) -1}  *100

% Difference = (ACVER / DC – 1) * 100  =   _____________