Helium Filling:

Needed when helium level drops below 20%.

Helium usage of He3 STM ~ 1 l/hr.

Note: Magnet is cold when resistance is higher than ~3.88 kOhm, corresponding to a Helium level > ~4%.

- Helium level meter on fast.

- Check weather extension piece of long transfer line is fixed.

- Put adaptor on dewar & close valve of dewar to return line to build up pressure.

- Slightly open needle valve of 1 K pot.

- Go up: Lower short transfer line a little bit to have outgoing helium gas.

- TIGHTEN ROPE & put “ear plug”

- Transfer line in vessel (keep ~30 cm above lowest point to avoid kinks) (Pressure max. 0.1- 0.2 bar).

- Connect both transfer lines

- Lower long transfer line

- Lower short transfer line, down to 1.7 mtr (when 50 % full) or 1.1 mtr (20 % full).

- Open valve of dewar to return line & start pumping on return line to reach pressure of 0.2 bar.

- Stop at 100 % full. Stop pump & depressurize dewar.

- Lift long transfer line as much as possible

- Lift short tube above head & tighten with rope & put “ear plug”

- Put level meter on slow or off.

NITROGEN FILLING:

- Pressurize dewar up to 0.5 bar

- Slowly open valve.